Telemetric Measurement of the Bone Healing with instrumented Implant – Comparison with CT-Investigations.

Majed Aljudaibi, Seide, Kowald, Karkaya, Weinrich, Jürgens, Faschigbauer

Berufsgenossenschaftliches Unfallkrankenhaus Hamburg

Clinical Question: The evaluation of bone healing occurs normally only radiologically. Often, for ex. by irregular nonunion, is a computed tomography (CT) as a "gold standard" required. Using modern microelectronics, an instrumented fixture was developed, in which transcutaneous Telemetry can determine the bending stress of the implant as a correlate of loss of elasticity in the healing bone. We investigated the hypothesis if with telemetric measurements; diagnosed stages of healing in CT, can be distinguished.

Method:

Femoral internal fixture systems (Tifix®) were instrumented with strain gauges and microcontroller-based mini telemetry electronics. A portable reading device enabled the wireless transcutaneous reading of the implant from a distance of 8 cm and the registry on a PC. Measurements were made at intervals of 2 weeks. By 27 patients with femoral shaft nonunion, a CT scan took place in addition after 12 weeks to classify the finding.

Results:

All nonunions consolidated, at a median of 21 weeks (minimum 12 weeks - maximum 37 weeks). Particularly interesting was that before a radiographic classification the elasticity in healing bone decreased significantly. The measured elasticity decreased to 50% at a median of 7 weeks (3.5 weeks - 13 weeks), to 30% (0% -56%) at 12 weeks and 10% after 19 weeks (14.4 weeks - 37 weeks). The measured values ​​differed between the CT stages of healing significantly (Kruskal-Wallis test, p = 0.0065).

Conclusion:

It was confirmed statistically that telemetric-mechanical measurements with instrumented implants enables the monitoring the bone healing. This can be done continuously and by avoiding the exposure to radiation.